## Slide 1
DISTRIBUTED EXPERIMENTS
LabOS
One experiment. Multiple PCs.
A connected measurement system.
Nonlinear Hall transport + bulk photovoltaic response
PC A: optical resources
SuperK + infrared spectroscopy
Central scientific agent
Planning + orchestration + evidence
PC B: sample and readout
attoDRY + electrical / optical control
A system-level view of hardware, references, timing, data and scientific decisions
LabOS | Cross-PC nonlinear Hall and BVPE measurement workflows
01/22
Notes: This presentation explains LabOS through a concrete two-PC instrument scenario. It is a capability and architecture presentation, not a report of measured sample results. The central idea is to organize optical excitation, cryogenic environment, electrical transport and photovoltage or photocurrent readout into reproducible stages with shared evidence.

## Slide 2
01 / SCIENTIFIC SCENARIO
Two response channels, one experimental context
Electrical and optical excitation share the same sample coordinates: T, B and gate voltage Vg.
Nonlinear Hall transport
Drive: lock-in sine output + series resistor
Readout: longitudinal Vxx and transverse Vxy
Axes: current, frequency, T, B and Vg
Controls: harmonics, return points,
repeatability
BVPE / photocurrent response
Drive: SuperK or a 730 nm laser
Reference: optical chopper at fchop
Axes: optical power, polarization, spectrum,
T/B/Vg
Controls: dark, power reference and repeat
scans
Shared sample
Readback coordinates
Controlled stages
Linked evidence
The research task defines the sequence; device ownership defines where each action executes.
LabOS | Cross-PC nonlinear Hall and BVPE measurement workflows
02/22
Notes: The electrical branch tests transverse second-order response under AC current drive. The optical branch measures chopped-light photocurrent or photovoltage while changing power, polarization and excitation spectrum. Both share sample identity, wiring configuration and environmental readbacks. A common context enables comparisons without assuming a common microscopic mechanism.

## Slide 3
02 / FIGURE 1 — SYSTEM ARCHITECTURE
An evidence-driven hierarchy across two control PCs
Scientific reasoning, deterministic orchestration and local instrument ownership are separate layers.
Researcher
Question, sample, limits
and physical interventions
Central scientific agent
Planning
Analysis
Troubleshooting
Evidence and memory
Papers / manuals / protocols
Code versions / previous runs
Deterministic execution coordinator
Schema + units
Limits + locks
Stage state
Command ledger
QC + hashes
PC A agent: optical worker
Persistent sessions / drivers / local watchdog
SuperK
Source + band control
CMD
STATE
IR spectrometer
Reference spectrum
TRIG
DATA
Command journal + local raw-data outbox
PC B agent: environment, transport and optics
Persistent sessions / drivers / local watchdog / artifact cache
attoDRY
T / B state
I/O
Transport x2
Vxx / Vxy
I/O
Optical chain
PD / laser
I/O
Gate + pol.
2400 / motor
I/O
BVPE lock-in + CLD1000 + power / polarization configuration
DOWN: approved commands UP: data, readbacks and faults LOCAL: session ownership and recovery state
Architecture presentation inspired by QUANTA Fig. 1 [1]; original LabOS redraw for this instrument scenario.
03/22
Notes: The operator supplies a scientific objective, sample constraints and physical interventions. A central scientific agent combines planning, analysis and troubleshooting with a versioned evidence store. Proposed stages pass through deterministic schema, units, limits, resource-lock and command-ledger checks. Persistent workers on PC A and PC B own device sessions, execute bounded actions and retain local evidence. Commands flow down; measurements, observed states and faults flow back up. The small device panels identify resources; later figures expand their actual signal paths.
Reference: Architecture presentation inspired by QUANTA Fig. 1 [1]; original LabOS redraw for this instrument scenario.

## Slide 4
03 / LOCAL WORKER ARCHITECTURE
A PC agent is a persistent instrument service
The central server schedules stages; the worker owns sessions, verifies actions and protects local evidence.
INSIDE EACH WINDOWS PC
Network client
Outbound poll / heartbeat
Persistent worker
Device ownership / dispatch
Driver adapters
Finite device capabilities
Instrument sessions
Vendor I/O and state
Command journal
IDs / outcomes / recovery
Local watchdog
Contact loss / stop latches
Raw-data cache
Atomic local artifact writes
Artifact outbox
Checksum / retry /
acknowledge
A network interruption does not erase a command, transfer ownership or discard a saved measurement.
LabOS | Cross-PC nonlinear Hall and BVPE measurement workflows
04/22
Notes: A local worker is more than a remote shell. It keeps instrument connections, translates finite capability calls through adapters, manages one owner per device and records each command identity. Status monitoring and the watchdog run independently of a long scientific conversation. Raw measurements are written locally before an artifact outbox synchronizes them with checksums. An ambiguous dispatched action is reconciled, not blindly replayed.

## Slide 5
04 / RESOURCE MAP
The complete two-PC instrument inventory
Device ownership stays local; instrument capabilities can participate in a shared multi-PC experiment.
Owner
Instrument / resource
Experiment role
PC A
SuperK supercontinuum source
Optical source and excitation-band configuration
PC A
Infrared spectrometer
Incident / reference spectrum and optical diagnostics
PC B
attoDRY
Temperature, magnetic field and sample-state readbacks
PC B
Transport lock-in x2
Master drive + Vxx; Vxy with selected harmonic
PC B
Chopper + BVPE lock-in
Optical modulation and referenced photoresponse
PC B
Photodiode (PD)
Optical-power monitoring and sample-plane calibration
PC B
2400 source-measure unit
Gate voltage Vg and gate leakage Ig
PC B
730 nm laser + CLD1000
Laser excitation and drive-current control
PC B
Motorized polarizer / quarter-wave plate
Linear-polarization angle / polarization state
Two transport lock-ins + one BVPE lock-in; one configured source path and one shared sample context.
LabOS | Cross-PC nonlinear Hall and BVPE measurement workflows
05/22
Notes: This inventory follows the supplied apparatus. PC A owns SuperK and the IR spectrometer. PC B owns attoDRY, two electrical transport lock-ins, a separate BVPE lock-in, a chopper, a power-monitoring photodiode, a 2400 gate source, the 730 nm laser and CLD1000 current controller, and a motorized polarization element. The electrical drive is a function of one transport lock-in, not a fourth lock-in.

## Slide 6
05 / FIGURE 2 — EXPERIMENT TOPOLOGY
One sample connects the optical and electrical chains
Control ownership follows PCs; light, voltage, current and physical references follow the apparatus.
PC A: source + spectrum
SuperK
Excitation band / spectrum
IR reference spectrum
PC B: laser excitation
730 nm laser
CLD1000 current control
Drive setting -> calibrated P
Polarizer / QWP
Chopper
PD monitor
Power calibration
attoDRY
Transport master
Drive f_I + Vxx(1f)
Transport follower
REF f_I -> Vxy(2f)
BVPE lock-in
REF fchop -> photoresponse
Chopper REF
fchop
2400: Vg / Ig
Legend: teal = optical / photoresponse; blue = transport readout; amber dashed = physical reference.
LabOS | Cross-PC nonlinear Hall and BVPE measurement workflows
06/22
Notes: The sample can receive light from PC A's SuperK or PC B's 730 nm source. A configured reference branch feeds the IR spectrometer; a power-monitoring branch supports sample-plane power calibration. PC B provides temperature, magnetic field, gate bias, electrical excitation and both transport and optical lock-in readout. Optical and electrical acquisitions are staged according to the selected wiring configuration. The cryostat drawing is an optical-access schematic, not a vendor CAD model.

## Slide 7
06 / CONTROL PLANE
Orchestration travels over HTTP(S), not signal cables
Windows agents initiate outbound connections; device commands are durable, scoped and attributable.
Central coordinator
Registry / reservation / stage ledger
PC A agent
Poll + heartbeat + artifact outbox
PC B agent
Poll + heartbeat + artifact outbox
APPROVED COMMAND
command_id / device / stage
operation / parameters / deadline
OBSERVED EVIDENCE
readbacks / timestamps / flags
artifact_id / checksum / outcome
Shared experiment identity
No SSH dependency: the network coordinates state; physical references establish lock-in phase.
LabOS | Cross-PC nonlinear Hall and BVPE measurement workflows
07/22
Notes: PCs poll the central endpoint, return status through heartbeats and synchronize measurement artifacts. Central-to-PC arrows describe command payloads returned to these clients, not a requirement for inbound SSH access. An instrument resource is reserved for a run before dispatch. Command identity, boot session and ownership protect against accidental duplication or stale sessions. Local raw data are acknowledged separately from the command outcome.

## Slide 8
07 / FIGURE 3 — TIMING AND SYNCHRONIZATION
A point barrier links conditions before acquisition
A shared point_id provides correspondence; acquisition intervals and physical references provide timing context.
Coordinator
PC A
PC B
Electrical ref.
Optical ref.
Issue point_id
Join READY states
Release acquisition
QC + archive
Set source / band / optics
Readback + READY
Reference spectrum
Set T / B / Vg; settle
Readback + READY
Lock-ins + PD
Barrier
f_I: master oscillator -> follower REF; select harmonic n
fchop: chopper controller -> BVPE lock-in REF
Store conditions AND time windows. Shared coordination is not a claim of zero-skew sampling.
LabOS | Cross-PC nonlinear Hall and BVPE measurement workflows
08/22
Notes: At each experimental point, optical preparation on PC A and environmental preparation on PC B can proceed independently. Each worker reports the requested point identity, actual readback and readiness evidence. The coordinator releases acquisition only after the required participants are ready. Each trace records its own start and end time. The electrical reference fI and optical chopper reference fchop remain physical signals, independent of network latency. Hardware triggering can be added where the experiment requires strict simultaneous sampling.

## Slide 9
08 / PREPARATION MANIFEST
Translate the apparatus into a reproducible run setup
Reusable device capabilities are combined with sample-specific wiring, calibration and experimental limits.
Record
Configuration content
Used by
Device ownership
PC A/B, resource IDs, connection and driver versions
Dispatch + locks
Sample and wiring
Sample ID, contacts, Vxx/Vxy mapping, gate, optical path
All stages
Electrical convention
Rs calibration, actual I, RMS/peak, f_I, harmonics, phase
Hall analysis
Optical convention
Source, excitation spectrum, PD calibration, polarization zero
BVPE analysis
Environment
T/B/Vg targets, readbacks and settling criteria
Point barrier
Action envelope
Approved drive, optical power, gate leakage and rates
Validation + QC
End-state policy
Excitation off; gate and environment follow defined policy
Run closure
A setup is both a resource map and an explicit statement of what each measurement means.
LabOS | Cross-PC nonlinear Hall and BVPE measurement workflows
09/22
Notes: The setup manifest describes physical ownership, instrument interfaces, wiring and the calibration conventions relevant to a campaign. It distinguishes a device limit from a sample-specific authorized envelope and from settings already supported by prior data. The user can reuse the device setup while changing the objective, scan coordinates, reference tests or closure policy for a new experiment.

## Slide 10
09 / FIGURE 4 — NONLINEAR HALL CIRCUIT
Two transport lock-ins, with explicit drive and sense paths
The master lock-in supplies AC excitation and Vxx(1f); the follower measures Vxy at a selected harmonic.
MASTER LOCK-IN
Internal oscillator: f_I
Vxx HI
Vxx LO
Drive return
Sine OUT
SAMPLE
I+
I-
Rs
I_rms = V_Rs,rms / Rs
FOLLOWER LOCK-IN
Differential Vxy input
External REF = f_I
Harmonic n = 2
Save X, Y and lock status
2400: Vg + Ig
Physical oscillator REF: f_I
Crossings without dots are not connections.
Vxx / Vxy are spatial channels. X / Y are phase components returned by EACH lock-in.
Circuit-level scenario based on the supplied apparatus; harmonic measurement rationale: [2,3].
10/22
Notes: The master transport lock-in both generates the sine wave and reads longitudinal voltage using its internal reference. Its oscillator reference is routed physically to the follower, which selects harmonic two for a transverse second-order measurement. The sample current is calibrated from the series-resistor voltage with an explicit RMS convention. The drive return is drawn separately from voltage sense leads. The transverse and longitudinal inputs may share a sample contact according to the selected differential wiring. The 2400 provides an independent gate voltage with leakage monitoring.
Reference: Circuit-level scenario based on the supplied apparatus; harmonic measurement rationale: [2,3].

## Slide 11
10 / HARMONIC ACQUISITION MODES
Change acquisition mode, preserve channel meaning
The same two transport instruments support a sequence of complementary harmonic and reference stages.
Stage
Master: Vxx
Follower: Vxy
Scientific purpose
A: transport baseline
1f, internal REF
1f, external f_I
Linear response and pickup baseline
B: second-order test
1f, internal REF
2f, external f_I
Vxy(2ω) versus actual Iω²
C: harmonic control
2f, internal REF
2f, external f_I
Longitudinal / transverse comparison
D: return reference
Chosen reference mode
Same reference settings
Repeatability and drift check
Preserve the full readout
X / Y, frequency, harmonic, gain, range and flags
Keep inference separate
Quadratic response is evidence, not a unique mechanism
LabOS | Cross-PC nonlinear Hall and BVPE measurement workflows
11/22
Notes: The first configuration reads Vxx at the fundamental and Vxy at the second harmonic. A subsequent baseline stage can select fundamental transverse voltage, and another stage can inspect longitudinal second-harmonic pickup. These are distinct acquisition stages for single-demodulator instruments; they are not presented as four simultaneous channels from two generic lock-ins. Every record includes the harmonic, phase reference, range and RMS convention.

## Slide 12
11 / HALL CAMPAIGN LOGIC
Turn a nonlinear signal into a controlled comparison
Outer coordinates: T × B × Vg. Inner tests: drive amplitude, frequency, harmonic and return references.
Prepare state
T / B / Vg readbacks
Dark baseline
Low drive + 1f
Drive sweep
Actual Iω + 2f
Return controls
Repeat / f / harmonics
Evidence review
Qualify + compare
Configured measurements
Current bracket and return point
Frequency / phase / harmonic controls
Environmental and gate-leakage traces
Repeated reference acquisitions
Evidence for the next decision
Readable and referenced signals?
Consistent response at reference settings?
Drive scaling preserved across controls?
Which comparison resolves ambiguity?
The agent proposes the next discriminating measurement within the configured experimental envelope.
LabOS | Cross-PC nonlinear Hall and BVPE measurement workflows
12/22
Notes: A campaign starts from a low-drive dark reference, then tests excitation-amplitude dependence and returns to a reference setting. Frequency and harmonic controls help separate a reproducible second-order response from thermal effects, contact nonlinearities and pickup. Field symmetry must be interpreted for the chosen physical model rather than imposed using a universal linear-Hall rule. The diagrams describe measurement logic, not measured sample curves.

## Slide 13
12 / FIGURE 5 — BVPE OPTICAL AND READOUT PATH
Optical drive, power and referenced readout
The 730 nm and SuperK paths converge on the same configured polarization and sample environment.
SuperK
Band / source configuration
SOURCE
730 nm laser
Selected laser excitation
LIGHT
CLD1000: I_LD
Chopper
Polarizer
QWP / motor
attoDRY
Optical-access sample
PD power monitor
Responsivity / transfer K
POWER
IR reference spectrum
BVPE LOCK-IN
Photovoltage input
or current-mode readout
Gain / range / X / Y
External REF = fchop
Chopper REF
fchop
2400: Vg / Ig
fchop is the optical-intensity envelope reference; it is not the optical carrier frequency or Hall 2f_I.
Optical / photocurrent scenario for LabOS; BVPE experimental context: Osterhoudt et al. [4].
13/22
Notes: SuperK provides a configured spectral excitation state, while the 730 nm diode branch is controlled through CLD1000 drive current. A selected beam path passes through the chopper and polarization element to an optical-access sample mount. The reference spectrum and PD power monitor characterize the optical condition; a calibration transfer links the monitor to power at the sample plane. The BVPE lock-in measures photovoltage directly or photocurrent through its configured current input or current preamplifier. Its physical reference comes from the chopper, not from the network.
Reference: Optical / photocurrent scenario for LabOS; BVPE experimental context: Osterhoudt et al. [4].

## Slide 14
13 / CALIBRATION AND NORMALIZATION
Describe the optical state at the sample plane
Drive settings, monitor readings and sample-plane power are related through an explicit calibration.
Source setting
I_LD or band setting
PD readback
Detector + electronics
Calibration K
Wavelength / path
Sample-plane P
Window / polarization
Normalized signal
Current or voltage / P
Power definition
Monitor responsivity
Calibration transfer K
Sample-plane reference
Modulation amplitude
convention
Polarization definition
Polarizer angle zero
QWP angle and retardance
Input polarization state
Motor position readback
Electrical definition
Voltage or current readout
Transimpedance gain, if used
X / Y phase and lock status
Source-drain bias convention
For current readout: Iph = Vread / G_TIA. Power normalization must state its calibration and units.
LabOS | Cross-PC nonlinear Hall and BVPE measurement workflows
14/22
Notes: CLD1000 sets diode drive current; it does not by itself measure optical power at the sample. The PD reading is converted through detector response, monitor electronics and an optical transfer calibration. This calibration can depend on wavelength and the configured optical path, including window transmission and polarization elements. A photocurrent channel also needs its amplifier gain and the modulation-amplitude convention. These definitions are stored with each acquisition rather than inferred after the experiment.

## Slide 15
14 / BVPE CAMPAIGN LOGIC
Combine spectrum, power, polarization and environment
PC A supplies optical resources; PC B supplies polarization, modulation, sample state and electrical readout.
Dark / reference
Background + spectrum
Power calibration
Source -> PD -> P
Polarization scan
Linear angle / QWP
State sweep
T / B / Vg / spectrum
Return reference
Dark / P / repeat
Branch
Configured scan
Associated evidence
730 nm excitation
Drive setting × power × polarization
PD, angle, gain and fchop readbacks
SuperK excitation
Excitation band / recorded incident spectrum
Reference spectrum and power calibration
Environmental sweep
T × B × Vg with selected optical state
Actual state, settling and gate leakage
Reference return
Dark / reference power / repeated polarization
Background, stability and repeatability
A chopped-light signal is an observation; mechanism assignments require the matching control experiments.
LabOS | Cross-PC nonlinear Hall and BVPE measurement workflows
15/22
Notes: A BVPE campaign includes dark and reference acquisitions, optical-power characterization, polarization scans and environmental sweeps. The 730 nm branch is useful for fixed-wavelength power and polarization measurements. The SuperK branch is configured as a selected excitation band or a measured broadband illumination state. The IR spectrometer operates within its configured response band; the power monitor and optical transfer calibration provide the power context. The current or voltage response can then be compared under common sample-state coordinates.

## Slide 16
15 / JOINT CAMPAIGN GRAPH
A shared coordinate links dark and illuminated stages
campaign_id links the research objective; stage/run IDs preserve branches; point IDs link conditions.
Prepare T / B / Vg
PC B: observed state
Dark Hall stage
PC B: drive + Vxx / Vxy
Optical reference
PC A + PC B: spectrum / PD
BVPE stage
PC B: fchop + photoresponse
Return references
Dark / power / repeat
Evidence review
Quality + controls + comparison
Next state / follow-up stage
sample + wiring + observed T/B/Vg + source state + time windows + calibration versions
Compare linked conditions and controls before inferring relationships between electrical and optical responses.
LabOS | Cross-PC nonlinear Hall and BVPE measurement workflows
16/22
Notes: A joint campaign can prepare an environmental state, acquire dark transport, characterize the optical condition, acquire BVPE and then repeat a dark or power reference. PC A and PC B participate only where their capabilities are needed. Source or wiring changes are explicit stage transitions and may include operator actions. The campaign graph keeps branches and return references distinct instead of flattening them into one undifferentiated curve.

## Slide 17
16 / DATA AND PROVENANCE
A measurement is more than a scalar value
Keep requested values, observed state, raw channels, calibration and quality decisions in separate fields.
Identity and context
campaign / run / stage / point
sample ID and wiring version
plan + setup fingerprints
requested T / B / Vg
observed T / B / Vg / Ig
acquisition start / end times
source + polarization
configuration
Raw measurement channels
Vxx: X, Y, n, f_I
Vxy: X, Y, n, f_I
photoresponse: X, Y, fchop
input mode / gain / range
PD monitor and calibrated
power
reference spectrum S(λ)
local file + artifact checksum
Qualification and interpretation
reference lock / overload flags
settling / freshness evidence
accepted / rejected / historical
calibration and analysis versions
reference / return comparisons
observations vs. hypotheses
closure state and next decision
Raw evidence is retained; derived analysis and interpretation remain attributable and revisable.
LabOS | Cross-PC nonlinear Hall and BVPE measurement workflows
17/22
Notes: A complete measurement record joins identity, requested setpoints and observed conditions without conflating them. The raw electrical and optical channels carry harmonic and reference definitions. Detector readings, spectrum and calibration identifiers explain the optical state. Quality assessment is separately versioned so that rejected data remain available for audit. Hashes link immutable plans, setup versions, artifacts and analysis products. The schema shown here is a capability-level record model rather than a claim that all fields share one instrument sampling clock.

## Slide 18
17 / EXECUTION STATE AND QUALITY
Explicit state transitions govern the next action
A completed command, a synchronized artifact and a scientifically usable reading are distinct outcomes.
PREPARE
Set and read back
READY
Participants + criteria
ACQUIRE
Local raw-data write
SYNCHRONIZE
Artifact + checksum
QUALIFY
Flags + evidence
ACCEPT -> analyze
Preserve raw + calibration + context
REJECT -> diagnose
Keep raw / flags; stop affected excitation
UNKNOWN -> reconcile
Verify owner / command / physical state
Closure is a recorded state, not just an acknowledgement
Excitation off and readback verified; gate / T / B follow the selected end-state policy.
The next proposed stage is validated again and respects operator gates and the configured action envelope.
LabOS | Cross-PC nonlinear Hall and BVPE measurement workflows
18/22
Notes: The coordinator moves a point through preparation, readiness, acquisition, synchronization and qualification. Rejected evidence is retained and can motivate a diagnostic plan. Unknown dispatch outcomes require ownership and state reconciliation instead of replay. Safe closure records observed output states; stopping electrical or optical excitation does not implicitly ramp a magnet or gate to zero. Physical rewiring or sample changes are explicit operator steps.

## Slide 19
18 / USER WORKFLOW
Start from the research objective, not a remote desktop
Device configuration is reusable; each experiment supplies a sample, variables, controls and operating bounds.
"Map dark nonlinear Hall versus T, B and Vg; then measure 730 nm power and
polarization dependence at matching states, with return references."
Describe objective
Sample + question
Prepare setup
Resources + wiring
Review stages
Axes + controls
Execute together
State + data + faults
Review evidence
Findings + next step
PC side: identity / instrument connections / drivers / local data / local stop policy
Experiment side: sample / wiring / allowed ranges / coordinate order / references / closure
The scientist focuses on design and interpretation; LabOS organizes execution and evidence.
LabOS | Cross-PC nonlinear Hall and BVPE measurement workflows
19/22
Notes: The operator can describe a scientific task in natural language and review the resulting stages before execution. The example asks for dark nonlinear Hall and polarization-dependent BVPE at matched environmental coordinates. The PC-side configuration identifies resources, connections, drivers, data paths and local policies. The experiment-side configuration identifies the sample, scientific question, wiring, coordinate sequence and controls. The interface provides live state, data-quality feedback and an attributable run review.

## Slide 20
19 / REUSE AND SCALE
Reuse local code. Compose shared capabilities.
An instrument capability is portable at the planning layer even when its vendor environment remains local.
Existing instrument code
Connect / identify / status
Set / read back / settle
Acquire / write raw data
Stop / verify output state
Bounded local capability
Typed inputs and units
Defined return data and flags
Timeout / cancellation
behavior
Exclusive device session
Reusable scientific stage
Environment preparation
Hall harmonic acquisition
Optical calibration / BVPE
Reference and diagnostic
tests
Central resource / capability graph
PC A: optics
PC B: transport
PC C: imaging
PC N: new resources
Scale by adding owned resources and verified capabilities, not by moving every instrument to one PC.
LabOS | Cross-PC nonlinear Hall and BVPE measurement workflows
20/22
Notes: Existing test scripts are valuable starting points for finite capability adapters: connect, read status, set, settle, acquire, save and stop with verification. Drivers keep manufacturer dependencies and physical connections local. The central resource graph can include more PCs while preserving resource locks, shared physical paths and local ownership. Code and protocols also become versioned evidence for planning, which is different from executing arbitrary retrieved code.

## Slide 21
LABOS CAPABILITY MAP
One coordinated experiment.
Multiple instruments and PCs.
Coordinate
Optical sources, sample state, transport and modulation
Preserve meaning
References, harmonics, power calibration and timestamps
Close the loop
Qualified observations, controls and a justified next stage
Nonlinear Hall + BVPE: a concrete example of composable, evidence-driven experimentation
LabOS | Cross-PC nonlinear Hall and BVPE measurement workflows
21/22
Notes: LabOS connects a distributed set of laboratory resources around a scientific objective. In this scenario, it joins two PCs, two excitation paths, environmental and gate control, harmonic transport, and chopped-light readout. Deterministic execution and local ownership support dependable operation, while shared identities, calibrations, timestamps and quality records support scientific interpretation. The central agent can then propose the next useful comparison from the evidence.

## Slide 22
APPENDIX / REFERENCES
Scientific context and diagram provenance
All architecture, apparatus, circuit, timing and state diagrams are original LabOS scenario drawings.
[1] Scientific-agent architecture
Rha et al., Autonomous Quantum Measurement with Agentic AI (2026).
https://arxiv.org/abs/2610.38850
[2] Nonlinear Hall theory
Sodemann and Fu, Physical Review Letters 115, 216806 (2015).
https://doi.org/10.1103/PhysRevLett.115.216806
[3] Nonlinear Hall measurement
Ma et al., Nature 565, 337-342 (2019).
https://doi.org/10.1038/s41586-018-0807-6
[4] Bulk photovoltaic response
Osterhoudt et al., Nature Materials 18, 471-475 (2019).
https://doi.org/10.1038/s41563-019-0297-4
Apparatus and interfaces: the supplied two-PC configuration. Figures explain system capabilities and workflows.
LabOS | Cross-PC nonlinear Hall and BVPE measurement workflows
22/22
Notes: The instrument inventory and drive method were supplied and confirmed by the user. QUANTA is the framework reference for separating scientific reasoning, deterministic execution and persistent hardware ownership. The nonlinear Hall papers motivate fundamental and second-harmonic readout and the need for controls. The BVPE paper provides optical-response context. These sources are not used as measured data from this apparatus. The cryostat and front-panel illustrations are conceptual, with roles and interfaces expressed through experiment configuration.