## Slide 1
LabOS | A Transport Experiment That Calibrates and Learns
Device communication → autonomous calibration → live feedback → physics-guided next experiment
A / DEVICE COMMUNICATION
B / MEASUREMENT AND FEEDBACK
C / SCIENTIFIC NEXT-STEP DESIGN
Commands down • evidence back
Central tools → bounded command
command_id / stage_id / device_id / deadline
PC-initiated HTTP(S)
Poll + heartbeat
Results + raw artifacts
PC i … PC N / local PC Agent
Device sessions + command journal + watchdog
Readbacks / flags / timestamps / local raw cache
Driver adapters / reviewed local code
VISA I/O or vendor SDK; SCPI only where supported
GPIB / USB / serial / TCP-IP
Lock-in master: sine OUT + Vxx(1f)
Lock-in follower: selected Vxy(2f)
attoDRY: T / B | 2400: Vg / Ig
Physical signal and reference cables
Sine OUT
Rs
Sample + return
Master f_I → follower REF; select harmonic
Network timing and physical phase reference
are separate parts of the experiment.
1 Define the transport campaign
Sample / wiring / T, B, Vg / current and frequency envelope
2 Autonomous calibration
Verify actual current, reference and phase
Choose range / filtering; repeat baseline checks
Reuse valid calibration or issue a new calibration_id
3 Prepare + acquire the next point
Join actual T/B/Vg readiness; settle after setting changes
Recheck calibration validity at the observed state
Save Vxx(1f), Vxy(2f), X/Y, readbacks and timestamps
4 Live qualification
Reference lock / overload / freshness / drift / leakage
Recoverable QC issue → recalibrate; hard fault → pause
5 Central feedback + scientific analysis
Fit amplitude / phase / drive dependence; compare controls
Propose a repeat, a new setting or a discriminating test
6 Validate update / close the campaign
Limits + locks + versioned update / recorded closure
Qualified evidence
Objective met / budget reached → archive + close
PLANNED • SERVER-LOCAL THEORY
Theory MCP / numerical jobs / versioned models
Receive a qualified-data snapshot
X/Y, I, f, T/B/Vg + calibration + snapshot IDs
Run models + estimate uncertainty
Transport / circuit / thermal-response hypotheses
Compare predictions with experiment
Residuals, validity range and competing explanations
Rank informative next measurements
Model discrimination / uncertainty / time / constraints
What changes after feedback?
Noisy but stable
Longer averaging / suitable filtering
Overload / reference drift
Correct range / reference; reacquire
Nonlinearity / possible heating
Lower drive; add frequency / return controls
Fault / uncertain command
Pause, local closure and state reconciliation
Every revision preserves raw evidence and passes
through the same execution authority.
LOCAL DEVICE LOOPS
Reference / readout / local fault response
No LLM in the timing-critical path
CENTRAL POINT / CHUNK FEEDBACK
Event-driven QC and bounded stage revisions
Dashboard follows state and qualified data
PLANNED ASYNCHRONOUS THEORY
Compute budget + uncertainty + snapshot checks
A local LLM is not a numerical physics solver
Transport workflow design; the theory/compute coupling is planned. No measured curves or material mechanism claims are implied.
Notes: This is a transport workflow design, complementing the general LabOS architecture. The left column separates the server-to-worker HTTP(S) link, local driver APIs and physical instrument interfaces. A master lock-in supplies sine excitation through a series resistor and reads Vxx at the fundamental; a follower reads Vxy at a selected harmonic using a physical reference. Dedicated calibration stages reuse the configured readout paths; no extra simultaneous measurement channel or automatic rewiring is assumed. Autonomous calibration adjusts bounded software settings, repeats references and issues a versioned calibration record. Raw X/Y and reference phase are retained, rather than auto-phasing an unknown second-harmonic signal to zero. The central loop qualifies each completed point or chunk, proposes a revision and validates a new stage version before dispatch. Recoverable quality problems can return to calibration; faults or uncertain commands require pause and reconciliation. The server-local theory/compute coupling is explicitly planned. It receives versioned qualified-data snapshots, compares transport, thermal and circuit hypotheses and ranks informative candidate measurements within the operating envelope. This is not a claim of deployed transport autonomy or measured material results.
Reference: Original LabOS workflow design; supporting primary sources are linked in the bilingual companion document.